## Test Matrix
FM Laser Testing Categories | Laser Campaign | Run | # of Pulses | Main Structure Temperature | Impact Location (1-5) | Azimuthal Location (1-3) | Impact Charge Range [pC] | Sensor Voltage | Parabolic Grid Voltage | Target Voltage | Detector Voltage | GSE or Flight CSAs Used | Duration | Notes
Room Temp. Detector Voltage Sweep | 1 |  | 10 | 20 | 3 | 1 | open | 2940 | 2868 | 2500 | 2200 | Flight | 0.25
| 1 |  | 10 | 20 | 3 | 1 | open | 2940 | 2868 | 2500 | 2400 | Flight
| 1 |  | 10 | 20 | 3 | 1 | open | 2940 | 2868 | 2500 | 2600 | Flight
| 1 |  | 10 | 20 | 3 | 1 | open | 2940 | 2868 | 2500 | 2000 | Flight
| 1 |  | 10 | 20 | 3 | 1 | open | 2940 | 2868 | 2500 | 1600 | Flight
Room Temp. Detector Charge Range | 1 |  | 200 | 20 | 3 | 1 | open | 2940 | 2868 | 2500 | TBD | Flight | 0.5
Flight Nom. Temp. Detector Charge Range | 1 |  | 200 | TBD | 3 | 1 | open | 2940 | 2868 | 2500 | TBD | Flight | 0.5 | IDEX's nominal temperature is TBD
[12 empty] | Total | 2 | days
General Notes:
Highlighted items are the parameters being varied
Non room temperature tests will have their thermal transistions occur overnight with testing occuring the following day.

## Nominal Configurations
HV Ramp Check
HV Channel | Voltage | Detector Resistance [Mohms] | In Parallel with 1G | Expected Current [uA] |  | Measured Current [uA]
Detector | 3006 | 162.812 | 140.0158 | 21.469 | HVPSDETCUR | 21.44
Parabolic | 3349 | 410 | 290.7801 | 11.5173
Target | 2509 | 50 | 47.619 | 52.689
|  |  | combined | 64.2063 | HPVSSENSPCUR | 67.74
Rejection | 368 | infinite | 1000 | 0.368 | HPVSSENSNCUR | 11.4
HV Nominal Values Check
HV Channel | Voltage | Detector Resistance [Mohms] | In Parallel with 1G | Expected Current [uA] |  | Measured Current [uA]
Detector | 2002 | 162.812 | 140.0158 | 14.2984 | HVPSDETCUR | 15.7
Parabolic | 2946.5 | 410 | 290.7801 | 10.1331
Target | 2504 | 50 | 47.619 | 52.584
|  |  | combined | 62.7171 | HVPSPSEN
Rejection | 198 | infinite | 1000 | 0.198 | HVPSNSEN
Noise Trigger for above parameters:
TOF HG max | 519 | 2200V | Threshold set to 570 = 10% over max
TOF HG max | 520 | 2000V | Threshold set to 530 = 2% over max
TOF HG max | 519 | 2400V | Threshold set to 530 = 2% over max

## 12-05-23 FM Laser Day 1
Test date December 5th, 2023
| Run | Temperature | Impact Location (see diagram) | Ref. Voltage | Target Voltage | Detector Voltage | GSE or Flight CSAs Used | AID USED | Number of Events | Laser Setting | Threshold | CSA | Gain
Laser Instrument Settings | 1 | 20 (RT) | 2 | 2868 | 2500 | 2200 | Flight | #8FF63550 | 36 | 4 | 570 | Target High Range (Low Gain) | 31936.65 | pC/V | 7984.1625
| 2 | 20 | 2 | 2868 | 2500 | 2000 | Flight | #8FF63550 | 22 | 3.5 | 570 | Target Low Range (High Gain) | 319.23 | pC/V | 23.9423 | 0.31325
| 3 | 20 | 2 | 2868 | 2500 | 2000 | Flight | #8FF63550 | 21 | 3.5 | 530 | Ion Grid | 1.499 | pC/V
| 4 | 20 | 2 | 2868 | 2500 | 2400 | Flight | #8FF63550 | 25 | 3.5 | 530
| 5 | 20 | 2 | 2868 | 2500 | 2400 | Flight | #8FF63550 | 37 | 4.5 | 530
| 6 | 20 | 2 | 2868 | 2500 | 2400 | Flight | #8FF63550 | 42 | 4 | 530
| 7 | 20 | 2 | 2868 | 2500 | 2400 | Flight | #8FF63550 | 54 | 3.5 | 530
Notes: | [12 empty] | Gold | 19.6499
First 3 and then 3 more at different laser focuses, then adjusted focus and got the remaining 30 hits. Focus was at 15.47mm | [12 empty] | Al | 7.2746
Second run, reduced laser setting to 3.5 and detector voltage to 2000V. | [12 empty] | Ti | 9.6995
Third run, reduced threshold to 530DN, only seeing peak at 20usecs after impact (gold?) | [12 empty] | Ni | 10.7262
Fourth run, increased detector voltage to 2400V
fifth run, increased laser to 4.5 and we saturated the target signal.
Sixth run, decrease laset to 4.0

## 12-06-23 FM Laser Day 2
Test date December 5th, 2023
| Run | Temperature Base Plate | Impact Location (see diagram) | Ref. Voltage | Target Voltage | Detector Voltage | GSE or Flight CSAs Used | Number of Events | Laser Setting | Threshold | CSA | Gain
Laser Instrument Settings | 1 | -33 | 2 | 2868 | 2500 | 2200 | Flight | 22 | 3.5 | 530 | Target High Range (Low Gain) | 31936.65 | pC/V | 7984.1625
| 2 | -33 | 2 | 2868 | 2500 | 2000 | Flight | 37 | 3.5 | 550 | Target Low Range (High Gain) | 319.23 | pC/V | 23.9423 | 0.31325
| 3 | -33 | 2 | 2868 | 2500 | 2000 | Flight | 70 | 3.5 | 550 | Ion Grid | 1.499 | pC/V
Notes: | [11 empty] | Gold | 19.6499
Run 1 starting at laser setting of 3.5 and threshold of 530. Moving laser spot with mirror, turning off LN2 plate attached to instrument, defocusing laser and then refocusing with translation stage. | [11 empty] | Al | 7.2746
Run 1 saw a lot of noise on the TOF which triggered an event. Increasing threshold to 550DN and will adjust focus for run 2 | [11 empty] | Ti | 9.6995
Run 2 threshold changed to 550DN | [11 empty] | Ni | 10.7262
Run 2 more small hits on TOF and ion grid, nothing on Target | [11 empty] | Na | 6.7142
Run3 adjusting focus, threshold same, adjusted focus back to position yesterday 16.48mm, saw everything from saturated Target LG to only Ion Grid and got some good TOFs | [11 empty] | K | 8.743
O