## Cover Page, Rev History
Laboratory For Atmospheric and Space Physics
LASP Engineering Division
University of Colorado
Boulder, Colorado
IMAP: IDEX - Interplanetary Dust Experiment
IDEX Calibration Test Matrix
LASP Document No. 172348revA
Approvers List
Prepared By |  | Scott Knappmiller, IDEX Lead Instrument Engineer
Approved By |  | Mihaly Horanyi, IDEX Principle Investigator
Approved By |  | Zoltan Sternovsky, IDEX CO Principle Investigator
Approved By |  | Kush Tyagi, IDEX Lead Systems Engineer
Approved By |  | Kristina Davis, IDEX Instrument Engineer
Configuration Management
Rev | Change Description |  | By
A | Initial Release |  | SK

## FM Dust Testing
FM Dust Accelerator Test Categories | Dust Campaign | Run | Dust Impacts | Dust Material | Impact Angle | Impact Location (1-5) | Azimuthal Location (1-3) | Velocity (5-50km/s) | Ref. Voltage | Sensor Voltage | Target Voltage | Detector Voltage | GSE or Flight CSAs Used | Duration [days] | Notes
Pre-Environmental Dust Testing
Detector Voltage | 1 | 1 | 60 | Iron | 0 | 5 | 1 | >20 | 2868 | 2939.7 | 2500 | 1700 | GSE | 1 | FM SN01 Detector Nominal Voltage for Gain of 1000 is 1720V
| 1 | 2 | 60 | Iron | 0 | 5 | 1 | >20 | 2868 | 2939.7 | 2500 | 1800 | GSE
| 1 | 3 | 60 | Iron | 0 | 5 | 1 | >20 | 2868 | 2939.7 | 2500 | 1900 | GSE
| 1 | 4 | 60 | Iron | 0 | 5 | 1 | >20 | 2868 | 2939.7 | 2500 | 2000 | GSE |  |  | Parabolic Grid | Sensor Voltage
| 1 | 5 | 60 | Iron | 0 | 5 | 1 | >20 | 2868 | 2939.7 | 2500 | 2100 | GSE |  |  | 2868 | 2939.7
Reflectron | 1 | 6 | 60 | Iron | 0 | 5 | 1 | >20 | 2918 | 2990.95 | 2500 | 1800 | GSE | 1 |  | 2918 | 2990.95
| 1 | 7 | 60 | Iron | 0 | 5 | 1 | >20 | 2968 | 3042.2 | 2500 | 1800 | GSE |  |  | 2818 | 2888.45
| 1 | 8 | 60 | Iron | 0 | 5 | 1 | >20 | 2918 | 2990.95 | 2500 | 1800 | GSE |  |  | 2768 | 2837.2
| 1 | 9 | 60 | Iron | 0 | 5 | 1 | >20 | 2868 | 2939.7 | 2500 | 1800 | GSE |  |  | 2718 | 2785.95
| 1 | 10 | 60 | Iron | 0 | 5 | 1 | >20 | 2818 | 2888.45 | 2500 | 1800 | GSE |  |  | 2668 | 2734.7
| 1 | 11 | 60 | Iron | 0 | 5 | 1 | >20 | 2768 | 2837.2 | 2500 | 1800 | GSE |  |  | 2618 | 2683.45
Impact Charge w/ Velocity | 1 | 12 | 60 | Iron | 0 | 5 | 1 | <5 | 2868 | 2939.7 | 2500 | 1800 | GSE | 1
| 1 | 13 | 60 | Iron | 0 | 5 | 1 | 5-8 | 2868 | 2939.7 | 2500 | 1800 | GSE
| 1 | 14 | 60 | Iron | 0 | 5 | 1 | 8-12 | 2868 | 2939.7 | 2500 | 1800 | GSE
| 1 | 15 | 60 | Iron | 0 | 5 | 1 | 12-20 | 2868 | 2939.7 | 2500 | 1800 | GSE
| 1 | 16 | 60 | Iron | 0 | 5 | 1 | 20-32 | 2868 | 2939.7 | 2500 | 1800 | GSE
| 1 | 17 | 60 | Iron | 0 | 5 | 1 | >32 | 2868 | 2939.7 | 2500 | 1800 | GSE |  | May increase # of dust impacts to decrease uncertainty in mass calculation
Impact Location | 1 | 18 | 60 | Iron | 0 | 1 | 1 | >20 | 2868 | 2939.7 | 2500 | 1800 | GSE | 1
| 1 | 19 | 60 | Iron | 0 | 2 | 1 | >20 | 2868 | 2939.7 | 2500 | 1800 | GSE
| 1 | 20 | 60 | Iron | 0 | 3 | 1 | >20 | 2868 | 2939.7 | 2500 | 1800 | GSE
| 1 | 21 | 60 | Iron | 0 | 4 | 1 | >20 | 2868 | 2939.7 | 2500 | 1800 | GSE
| 1 | 22 | 60 | Iron | 0 | 6 | 1 | >20 | 2868 | 2939.7 | 2500 | 1800 | GSE
Trigger Verification | 1 | 23 | 60 | Iron | 0 | 5 | 1 | >20 | 2868 | 2939.7 | 2500 | 1800 | GSE |  | Trigger method, could be moved to Post-Env test
Dust Material / Impact Charge w/ Velocity | 1 | 24 | 60 | Olivine | 0 | 5 | 1 | <5 | 2868 | 2939.7 | 2500 | 1800 | GSE | 3 | Repeat velocity measurements, dust source change takes a least a day
| 1 | 25 | 60 | Olivine | 0 | 5 | 1 | 5-8 | 2868 | 2939.7 | 2500 | 1800 | GSE
| 1 | 26 | 60 | Olivine | 0 | 5 | 1 | 8-12 | 2868 | 2939.7 | 2500 | 1800 | GSE
| 1 | 27 | 60 | Olivine | 0 | 5 | 1 | 12-20 | 2868 | 2939.7 | 2500 | 1800 | GSE
| 1 | 28 | 60 | Olivine | 0 | 5 | 1 | 20-32 | 2868 | 2939.7 | 2500 | 1800 | GSE
| 1 | 29 | 60 | Olivine | 0 | 5 | 1 | >32 | 2868 | 2939.7 | 2500 | 1800 | GSE |  | May increase # of dust impacts to decrease uncertainty in mass calculation
Impact Charge w/ Velocity w/ Flight CSAs | 1 | 30 | 60 | Olivine | 0 | 5 | 1 | <5 | 2868 | 2939.7 | 2500 | 1800 | Flight | 2 | Vacuum Break to switch GSE CSAs to Flight CSAs
| 1 | 31 | 60 | Olivine | 0 | 5 | 1 | 5-8 | 2868 | 2939.7 | 2500 | 1800 | Flight
| 1 | 32 | 60 | Olivine | 0 | 5 | 1 | 8-12 | 2868 | 2939.7 | 2500 | 1800 | Flight
| 1 | 33 | 60 | Olivine | 0 | 5 | 1 | 12-20 | 2868 | 2939.7 | 2500 | 1800 | Flight
| 1 | 34 | 60 | Olivine | 0 | 5 | 1 | 20-32 | 2868 | 2939.7 | 2500 | 1800 | Flight
| 1 | 35 | 60 | Olivine | 0 | 5 | 1 | >32 | 2868 | 2939.7 | 2500 | 1800 | Flight |  | May increase # of dust impacts to decrease uncertainty in mass calculation
[13 empty] | Total | 9 | days
Post-Environmental Dust Testing
Baseline Test | 2 | 1 | 60 | Olivine | 0 | 5 | 1 | open | 2868 | 2939.7 | 2500 | 2400 | GSE | 1
Voltage Sweep | 2 | 2 | 60 | Olivine | 0 | 5 | 1 | open | 2868 | 2939.7 | 2500 | if neccesarry | GSE
Impact Charge w/ Velocity | 2 | 4 | 60 | Olivine | 0 | 5 | 1 | 3-8 | 2868 | 2939.7 | 2500 | TBD | GSE | 1
| 2 | 5 | 60 | Olivine | 0 | 5 | 1 | 8-12 | 2868 | 2939.7 | 2500 | TBD | GSE
| 2 | 6 | 60 | Olivine | 0 | 5 | 1 | 12-20 | 2868 | 2939.7 | 2500 | TBD | GSE
| 2 | 7 | 60 | Olivine | 0 | 5 | 1 | 20-32 | 2868 | 2939.7 | 2500 | TBD | GSE
| 2 | 8 | 60 | Olivine | 0 | 5 | 1 | >32 | 2868 | 2939.7 | 2500 | TBD | GSE |  | May increase # of dust impacts to decrease uncertainty in mass calculation
Impact Location | 2 | 9 | 60 | Olivine | 0 | 4 | 1 | >20 | 2868 | 2939.7 | 2500 | TBD | GSE | 1
| 2 | 10 | 60 | Olivine | 0 | 6 | 1 | >20 | 2868 | 2939.7 | 2500 | TBD | GSE
Optional | 2 | 10 | 60 | Olivine | 0 | 5 | 1 | <2 | 2868 | 2940 | 2500 | TBD | GSE |  | Look for heavy Fe particles
Aluminum
| 2 | 11 | 60 | Aluminum | 0 | 5 | 1 | open | 2868 | 2939.7 | 2500 | TBD | Flight | 1
[13 empty] | Total | 4 | days
General Notes:
Highlighted items are the parameters being varied

## Olivine 12-3 thru 12-15
FM Dust Accelerator Test Categories | Date | Dust Run (Johns #) | # Sent Through from Accel | # of Hits achieved | Threshold | Dust Material | Impact Location (1-6) | Azimuthal Location (1-3) | Velocity (5-50km/s) | Ref. Voltage | Target Voltage | Detector Voltage | GSE or Flight CSAs Used | Rejection voltage | Notes
Open baseline | 2023-12-13 | 0 | 217 | 103 | 540 | Olivine | 5 | 1 | 2-90 | 2940 | 2500 | 2400 | GSE | 200 | No QD data on scope for the first 8 events (trigger was set to c1 and not ext trigger, then signals were inverted)
vel > 12 | 2023-12-13 | 1 |  | 27 | 540 | Olivine | 5 | 1 | 12-90 | 2940 | 2500 | 2400 | GSE | 200 | First 2 events are probably the gate valve opening
VbDet2500 | 2023-12-13 | 2 |  | 32 | 580 | Olivine | 5 | 1 | 12-90 | 2940 | 2500 | 2500 | GSE | 200 | First QD event is gate valve opening, event 15 is gate valve closing as John reconditions dust source
Detector Voltage 2500, Velocity gt 12km/sec | 2023-12-14 | 0 | 329 | 104 | 580 | Olivine | 5 | 1 | 12-90km/sec | 2940 | 2500 | 2500 | GSE | 200 | lowered velocity to 8km/sec at ~10:30am
Changed position to 4 | 2023-12-14 | 1 | 274 | 102 | 580 | Olivine | 4 | 1 | 8-90km/sec | 2940 | 2500 | 2500 | GSE | 200 | changed position to 4, moved 22.5 rotations
Notes from 12/14: high speed dust not flying very frequently. We still have microphonics from getter plate.
Velocity Scan at position 4 | 2023-12-15 | 0 | 1761 | 857 | 580 | Olivine | 4 | 1 | 8-90km/sec | 2940 | 2500 | 2500 | GSE | 200 | John is pushing the source and we are getting a lot of hits!
The scope got 849 hits

## Al 12-18 thru 12-20
FM Dust Accelerator Test Categories | Date | Dust Run (Johns #) | # Sent Through from Accel | # of Hits achieved | Threshold (DN) | Dust Material | Impact Location (1-6) | Azimuthal Location (1-3) | Velocity (5-50km/s) | Ref. Voltage | Target Voltage | Detector Voltage | Rejection voltage | GSE or Flight CSAs Used | Notes
run0 Al vel 3-90 det 2500 pos5 | 2023-12-18 | 0 | 87 | 84 | 580 | Aluminum | 5 | 1 | 3-90 | 2868 | 2939.7 | 2500 | -200 | GSE | potentially have an error in the recording of QD data on the IMPACT server. So we may not have any matching data for QD signals from the IMACT server metadata to compare to scope data | noise capture 8FF35217
run1 | 2023-12-18 | 1 |  | 120 | 580 | Aluminum | 5 | 1 | 3-90 | 2868 | 2939.7 | 2500 | -200 | GSE | aid 8ff3521c
run2 | 2023-12-18 | 2 |  | 71 | 580 | Aluminum | 5 | 1 | 10-90 | 2868 | 2939.7 | 2500 | -200 | GSE | aid 8ff3521d
run3 | 2023-12-18 | 3 |  | 112 | 580 | Aluminum | 5 | 1 | 10-91 | 2869 | 2940.725 | 2500 | -200 | GSE
run4 | 2023-12-18 | 4 |  | 90 | 580 | Aluminum | 5 | 1 | 10-92 | 2870 | 2941.75 | 2500 | -200 | GSE
run0 | 2023-12-21 | 0 |  | 96 | 580 | Al | 5 | 1 | 10-90 | 2870 | 2940 | 2500 | 2500 | flight
run1 | 2023-12-21 | 1 | 174 | 112 | 580 | Al | 5 | 1 | 12-90 | 2870 | 2940 | 2500 | 2500 | flight | The number of scope hits does not match the number of events form the instrument. We may be +/-1 due to quick dust hit on the scope, but tht won't account for greater mismatches.
run2 | 2023-12-21 | 2 | 175 | 123 | 580 | Al | 5 | 1 | 10-91 | 2870 | 2940 | 2500 | 2500 | flight | Scope ended at 118 hits, but collected first 3 at the same rate. It remains unclear why they end with different counts
run3 | 2023-12-21 | 3 |  | 110 | 580 | Al | 5 | 1 | 12-91 | 2870 | 2940 | 2500 | 2500 | flight | Saw at least one double hit on the scope, perhaps double hits explain the count misatch that occur over the course of an acquire
run4 | 2023-12-21 | 4 |  | 88 | 580 | Al | 5 | 1 | 10-92 | 2870 | 2940 | 2500 | 2500 | flight

## Position Diagram
2.245
2.245